# Interpretation of Tapestation Traces

Genomics Research Center

# Introduction to Tapestation Assays

Agilent Tapestation Instrumentation

# Available Tapestation Assays

|  | D1000 | High Sensitivity D1000 | D5000 | High Sensitivity D5000 | Genomic DNA |
| --- | --- | --- | --- | --- | --- |
| Volume | 1 uL | 2 uL | 1 uL | 2 uL | 1 uL |
| Size Range | 35 - 1,000 bp | 35 - 1,000 bp | 100 - 5,000 bp | 100 - 5,000 bp | 200 to >60,000 bp |
| Concentration | 0.1 - 50 ng | 0.01 - 1 ng | 0.1 - 50 ng | 0.01 - 1 ng | 10 - 100 ng |

The GRC will determine the most appropriate Tapestation assay for your sample type based on:

Expected Size (D1000 vs D5000)

DNA – D1000 + HS D1000

|  | D1000 | High Sensitivity D1000 |
| --- | --- | --- |
| Volume | 1 uL | 2 uL |
| Size Range | 35 - 1,000 bp | 35 - 1,000 bp |
| Concentration | 5 - 50 ng | 0.01 - 1 ng |

Ladder – A ladder is included with each run. The ladder contains two internal standards (Lower and Upper Markers) to align the ladder data with samples to determine sizing. The concentration of the upper marker is known and used to help determine sample concentration*.

D1000 Ladder	HSD1000 Ladder

*The GRC does not use Tapestation to determine concentration. Tapestation is used for quality assessment and

DNA – D1000 + HS D1000

Example Sample Trace – Lower and upper markers are included with each sample and are used to align the sample with the ladder. A region (Region 1) has been set and is indicated by the vertical blue lines on either side of the library trace. This is manually applied to each sample based on the library profile. This calculates the size of the library within this region and is used for normalization prior to sequencing.

Region 1

Region 1

D1000 Sample Trace	HSD1000 Sample Trace

DNA – D5000 + HS D5000

|  | D5000 | High Sensitivity D5000 |
| --- | --- | --- |
| Volume | 1 uL | 2 uL |
| Size Range | 100 - 5,000 bp | 100 - 5,000 bp |
| Concentration | 5 - 50 ng | 0.01 - 1 ng |

Ladder – A ladder is included with each run. The ladder contains two internal standards (Lower and Upper Markers) to align the ladder data with samples to determine sizing. The concentration of the upper marker is known and used to help determine sample concentration*.

D5000 Ladder	HSD5000 Ladder

*The GRC does not use Tapestation to determine concentration. Tapestation is used for quality assessment and

DNA – D5000 + HS D5000

Example Sample Trace – Lower and upper markers are included with each sample and are used to align the sample with the ladder. A region (Region 1) has been set and is indicated by the vertical blue lines on either side of the library trace. This is manually applied to each sample based on the library profile. This calculates the size of the library within this region and is used for normalization prior to sequencing.

Region 1*

Region 1

D5000 Sample Trace	HSD5000 Sample Trace

*Region 1 is often placed up to 1500 bp even if the library trace extends beyond this limit. In our experience, library fragments larger than 1500 bp do not cluster well and are not considered when determining library size for sequencing normalization.

# Unexpected Library Peaks

If peaks are present outside the expected library range, it could be due to primer dimers or adapter dimers. Further purification should be performed to remove these peaks prior to next generation sequencing.

Adapter

Dimer

# High Sensitivity Genomic DNA

Ladder – A ladder is included with each run. The ladder contains a lower marker (LM) to align the ladder data with samples to determine sizing, concentration, and DIN (DNA Integrity Number)

Example Sample Trace – Below is an example of a genomic DNA sample. A Lower Marker (LM) is included with each sample to align to the ladder data.

DIN: DNA Integrity Number: Used to assess genomic DNA degradation.

DIN is on a scale of 1-10 with 1 indicating severe degradation.

DIN: 7.4	DIN: 9.2

*Tapestation software does not allow a region beyond 60,000 bp. The Tapestation is not used for quantification

Contact Us

Genomics Research Center

Wilmot Cancer Center

601 Elmwood Ave., Box 704

Rochester, NY 14642

Phone: (585) 275-7250

Website

Email: URGenomics@urmc.Rochester.edu